TGL18台式高速冷冻离心机特点:

Model: TGL18 Table top high speed refrigerated centrifuge

Features:

Brushless frequency motor for model TGL18 which in great torque, free maintenance, no powder pollution, quick in speed up and down.

The flexible axle driven system which drive the rotor directly, smooth in operation, low noise and small vibration.

Import compressors fluorine free, double cycle cooling, cold and hot alternating easily, free environment pollution and precise in temperature control.

There are16 kinds of rotors for your choice.

Micro computer control system, digital display the temperature、time and speed. There are 10 kinds of program and 10 kinds of acceleration and deceleration for your choice.

Automatically electric lid lock, super speed、over temperature protection and imbalance protection.

The centrifuge body is made of high-quality steel, safe and reliable.

仪器特点：

。本仪器采用双核处理器，实现高低速兼容、大小容量兼容，是根据目前国际生物研究的实际要求所研发的最隹产品。

。采用无氟压缩机组，双循环制冷，冷热交换能力强，无环境污染，控温精准。

。ＴＧＬ18采用无碳刷变频电机、免维护，大力矩，升降速快。

。均采用微机处理器精准控制，数字显示转速、温度、时间等参数，按键编程，切换显示运行参数及RCF值。

。可存储和调用10套程序,10种升降速率供选择。

。门盖独立电机伺服，减轻操作人员劳动强度，超速、超温自动保护，不平衡保护，机身采用优质钢材结构；安全可靠，使用倍感方便。

。采用本公司特有的弹簧锥套连接转头与主轴，装卸转头快捷简单无方向性，安全可靠，使用倍感方便。

。配备多种转头供选择，并可根据试验要求设计各种适配器，一机多用。

。三级减振，离心效果达到最佳。

技术参数：

| 最高转速max speed： | 21000r/min | 最大离心力max RCF： | 29670×g |
| --- | --- | --- | --- |
| 最大容量max volume： | 4×750ml | 温控范围temperature range： | -20°C～40°C |
| 定时范围timer | 1～99h/59min | 噪 音noise： | ≤58dBA |
| 外型尺寸Dimension： | 688×715×430 | 净 重net weight： | 125KG |
| 控速精度Speed Accuracy： | ±50r/min | 电 源power supply： | AC 220V 50HZ 15A |
| 控温精度Temperature Accuracy： | ±1°C |  |  |

注册证号：湘长食药监械（准）字2007第1410045号

TGL18系列台式高速冷冻离心机适配角转头

Rotors for TGL18MC/TGL18C/TGL18M/TGL18

| 编号 | 转子类型 | 最高转速r/min | 容量ml | 最大离心力×g |
| --- | --- | --- | --- | --- |
| No30801 | 角转子 | 21000 | 12×1.5/2ml | 28030 |
| No30802 | 角转子 | 15000 | 40×0.5ml | 22920 |
| No30803 | 角转子 | 17000 | 24×1.5/2ml | 26460 |
| No30404 | 角转子 | 13500 | 30×1.5/2ml | 19340 |
| No30805 | 角转子 | 16000 | 16×5ml | 22020 |
| No30806 | 角转子 | 16000 | 12×7ml | 21380 |
| No30407 | 角转子 | 10000 | 12×15ml | 11840 |
| No30808 | 角转子 | 15000 | 12×10ml | 22680 |
| No30809 | 角转子 | 15000 | 8×20ml | 22680 |
| No30810 | 角转子 | 14000 | 6×30ml | 19060 |
| No30811 | 角转子 | 12000 | 6×50ml | 16050 |
| No30812 | 角转子 | 12000 | 6×70ml | 15570 |
| No30813 | 角转子 | 12000 | 4×100ml | 14850 |
| No30814 | 角转子 | 10000 | 6×100ml | 11380 |
| No30815 | 角转子 | 16000 | 6×10ml | 21500 |
| No30816 | 角转子 | 18000 | 30×0.5ml | 26660 |
| No30627 | 角转子 | 5000 | 30×15ml | 3830 |
| No30642 | 角转子 | 4000 | 24×50ml | 2970 |
| No30643 | 角转子 | 4000 | 12×100ml | 2970 |
| No30844 | 角转子 | 13000 | 48×1.5/2ml | 17930 |
| No30435 | 垂直转子 | 16000 | 16×5ml | 16540 |
| No30480 | 水平转子 | 15000 | 4×5ml | 19920 |
| No31377 | 水平转子 （多管架） | 5000 | 适配下列管托、管挂 |  |
|  |  |  | 4×50ml | 4730 |
|  |  |  | 4×100ml | 4730 |
| No31378 | 水平转子 （多管架） | 4000 | 适配下列管托、管挂 |  |
|  |  |  | 4×2×100ml | 3020 |
|  |  |  | 4×2×50ml | 3200 |
|  |  |  | 4×4×15ml | 3200 |
|  |  |  | 4×6×15ml | 3200 |
|  |  |  | 4×8×15ml | 3200 |
|  |  |  | 4×12×7ml | 2880 |
|  |  |  | 4×8×10ml | 2790 |
|  |  |  | 4×6×10ml | 2790 |
|  |  |  | 4×4×10ml | 2790 |
| No30695 | 水平转子(圆杯) | 4000 | 2×750ml(500ml试瓶) | 3380 |
| No30696 | 水平转子(圆杯) | 4000 | 4×750ml(500ml试瓶) | 3380 |
|  | 适配40301—40308适配器 |  |  |  |
|  | 水平转子(椭圆杯) | 4000 | 4×500ml | 3380 |
| No30592 | 水平转子 | 4000 | 4×250ml | 2920 |
|  |  | 适配40401—40408适配器 |  |  |
| No30679 | 水平转子(方杯) | 4000 | 4×500ml | 3310 |
|  |  | 适配40401—40408适配器 |  |  |
| No30589 | 水平转子 | 4000 | 4×22×5ml（真空管） | 2840 |
|  |  |  | 4×31×5ml（放免管） |  |
| No30590 | 水平转子 | 4000 | 4×22×7ml（真空管） | 3140 |
|  |  |  | 4×31×7ml（放免管） |  |
| No30691 | 水平转子（酶标板） | 4000 | 2×4×96孔 | 2490 |
| No30694 | 水平转子（酶标板） | 4000 | 4×4×96孔 | 2840 |
| No30676 | 酶标板转子 | 4000 | 2×3×48孔 | 2300 |